## Slide 1: verifying prototype Performance—cdr requirements
Dale E. Gary
Professor, Physics, Center for Solar-Terrestrial Research
New Jersey Institute of Technology

## Slide 2: outline
Open discussion about Critical Design Review requirements
CDR Decisions
Schedule for CDR
Long-lead Items
Questions
  Do we wait on all aspects of build-out until formal CDR is done, or give go-ahead on subsystems as they pass their criteria? Given the passage of time, this will set how agressive we must be on completing the testing.
  What aspects of performance have to be “passed” before go-ahead?
  What are the quantitative parameters that are to be checked and what are the specs?
  What subsystems are irrelevant for CDR? (e.g. antennas are already in hand, cross-dipole feeds are not open for reselection, etc.)

## Slide 3: CDR Decisions
There are several components of the prototype that are one-off items—no CDR decision needed, although they do need to be evaluated:
  Timing distribution module, LO distribution module, Correlator clock module, ACC, Network (also antennas are already in hand)
Build-out involves the following, which are subject to CDR considerations:
  8 more Tecom feeds (unlikely to be affected by CDR)
  10 more front-ends and components
  10 more auxiliary boxes and components
  10 more cRIOs (unlikely to be affected by CDR)
  10 more optical links
  10 more downconverter modules and components
  1 additional Hittite synthesizer
  6 more ROACH-2s and enclosures
  Arista switch and SFP+ modules and cables
  2 cryo-receivers and feed horns for 27-m
  2 compressors, coolers, for 27-m
  2 feed-positioning apparatuses (apparati?)

## Slide 4: Previously presented schedule =>Earliest CDR is feb 2013
2012 Oct-Nov:
  Dale & Gelu get ACC/cRIO control of antennas working + rudamentary schedule and stateframe.
  Wes sends hardware as it is completed to OVRO, where Kjell and Dan set it up with phone/email support.
  Toward end of period, correlator is shipped and set up with phone/email support.
2012 Early Dec:
  Wes, Dale & Gelu get front end installed and under control. Fix any problems with rest of system (may require Nimish).
  Initial attempt to determine feed best focus.
2012 Late Dec(?):
  Dale, Gelu, Nimish, Gordon, Jim set up DPP, attempt to get first fringes
2013 Jan:
  Dale, Gelu, Nimish, Gordon, Jim, Stephen attempt to determine baseline corrections.
  Initial attempt to get calibrated data—evaluation of phase and amplitude stability

## Slide 5: Suggest planning cdr for Mar 2013
Is that too late for proceeding with parts ordering for analog system? Probably, so how can we move that up?
  Kjell and Dan set up control-room hardware in Nov. 2012, Wes comes to OVRO in Dec. 2012, first observations in Jan. 2013.
  Probably no quantitative measurements will be available of power levels, etc., until Wes is at OVRO. Can we sign off on some of the long-lead items then, so that they can be ordered by Jan. 2013?
Items in BOM with 12wk or more lead time:
  Tecom feed (FE: 21 weeks)
  Optical Links (FE: 12-16 weeks)
  TE Cooler module (FE: 16 weeks)
  PWM Controller (Aux: 16 weeks)
  Teledyne + Q Microwave Amplifiers (DC: 12 weeks)
  ATM Power Splitter (DC: 12 weeks)
Other than these items, a 17-24 March 2013 CDR should be fine.

## Slide 6: Long-lead item check-off
To permit ordering of long-lead items in Jan 2013, we will need to fully check out the optical links and TE coolers of the front end.
  Some checkout (levels, etc.) will be done when Wes is at OVRO in Dec.
  Additional check of thermal and phase stability requires tests in Jan. 2013, so earlier the better.
Other long-lead items in AUX and DC can probably be checked off during Wes’ visit to OVRO in Dec. 2012.
With ordering of long-lead items in Jan., and CDR in March, parts should be in hand on the June timeframe—is this sufficient for build-out by, say, August, installation in Sep 2013? ALL money must be spent by 2013 Sep 30!

## Slide 7: Quantitative specs to meet
From Wes’s Cascaded Noise/Gain spreadsheet:
  Check power levels at relevant points in the system (power meter)—goal: match spreadsheet
  Check bandpass shape (spectrum analyzer, correlator output)—goal?
  Check power slope at relevant points (spectrum analyzer)—goal at IF output: -1 dB/BW
  Check ripple at relevant points (spectrum analyzer, correlator output)—goal: 4 dB
  Check system temperature (observations on calibrator, solar increment)—goal: 570 K
  Check IMD (how?)—goal: -10.2 dBc
  Check linearity (gain calibration)—goal: 1% in range ±3 dB around “sweet spot”
Temperature stability
  FE inside temperature—goal: within ?° with swing of ?° ambient.
  Power variation over temperature (noise diode increment over temperature range)—goal: 1% = 0.04 dB
  Amplitude and phase stability of temperature (calibrator observations)—goal: 1%, 1°
Antenna specs
  RMS pointing residuals <0.03 degree both optical star (all sky) and on Sun (radio, single Dec), repeatable over a full-day track on the Sun
  H, V beams match theoretical to within 5% to -10 dB (width, shape, taking into account dish illumination)
  Frequency-dependent pointing center stable to repeated measurement to within SNR expectation.
  Dishes track source (indicate lock) >99% of the time in 20 mph winds

## Slide 8: Quantitative specs to meet (cont’d)
Polarization
  X, Y cross-talk better than -12 dB average, -8 dB worst case all frequencies (how to measure?)
  X, Y corrected cross-talk better than -20 dB average, -15 dB worst case
  X, Y to R, L conversion verified, with isolation -20 dB average, -15 dB worst case
Control System Timing
  Band switching complete in <1 ms, independent of step.
  Attenuation control complete in <1 ms, independent of step.
  All stateframe subsystem timestamps identical to 100 ms, stable to 5 ms.
Correlator Timing
  Accumulation start and end fixed relative to 1pps tick within 2 spectral samples (~14 ms)
  Number of accumulations unvarying.
  Coarse delays adjusted on 1 pps tick, after last accumulation window.
What other critical specs do we have?